TO CHANGE/CLEAN OXYGEN SOURCE CHECKLIST

___	Set HV, Arc, and Coil to 0 in Sources window

___	Close O bottle, green knob, & valve on duo

___	Turn Freon pump off & open Freon tank for pouring back in

___	Flip off HV button on keyboard & 2 electronics switches in back of cabinet

___	Disassemble exterior Duo housing:

___	Remove 4 gray plates: Top, 2 sides, and Z electrode knob on front

___	Remove 2 Teflon sheets

___	Disconnect 4 electrical wires on chassis: Arc, coil, & 2 HV (anode & chathoed)

___	Disconnect two Freon lines

___	Unscrew nuts (4?) holding Duo chassis onto ceramic housing

___	VERY GENTLY remove needle valve: CAREFUL not to kink the O line

___	Remove Duo assembly:

___	Lift UP and keep horizontal

___	Carry to Freon tank and pour excess back in

___	Place on clean tray, undo pressure valve clamp, and put it all aside

___	Remove anode, Z-electrode, & cathode from assembly:

___	Using a slotted tool, unscrew anode counter-clockwise

___	Using special gripping tool, unscrew Z-electrode

___	Using an allen wrench, unscrew screw at bottom of cathode

___	Blow out all dirt from both ends in Duo chassis with compressed N

___	Replace anode, Z-electrode, & cathode with CLEAN, NEW (inkenel) parts:

___	Screw cathode on VERY TIGHT

___	Screw Z-electrode on

___	Screw anode back on to point of some resistance

___	Unscrew anode 2 complete revolutions counter-clockwise (correct height)

___	Reverse the assembly instructions:

___	Put pressure valve back on & place chassis onto instrument

___	VERY CAREFULLY attach needle valve

___	Screw chassis back onto ceramic, attach Freon lines, & connect 4 wires

___	Reassemble all external parts: Teflon sheets, Duo housing, & Z-electrode

___	Restart Freon pump & turn all electronics back on